## Slide 1: Multi-Frequency Resonator Development
Jason W. Sidabras
D. Suter, E.J. Reijerse, W. Lubitz

## Slide 2
Finite-Element Modeling
Ansys HFSS solves the full-wave Maxwell’s Equations for a given geometry and boundary condition
Build and Test on the Bench
Characterize the resonators and compare to simulation
Measure coupling, frequency, and Q-value
Experiments with Standard Samples
Further characterize the resonators using EPR
Sample handling
Magnetic field magnitude
Electric field magnitude
Resonator Introduction
Rectangular TE102
Notes: ----- Meeting Notes (3/30/17 16:58) -----
title methods

## Slide 3
Feher, Sensitivity Considerations in Microwave Paramagnetic Resonance Absorption Techniques, 36(2). 449,1957
Hyde, Froncisz, Advanced EPR: Applications in Biology and Biochemistry, Elsevier, 1989, Ch. 7: Loop-gap resonators, 277
Based on a critically coupled resonator and reflection bridge configuration
Resonator Characteristics
Cylindrical TE011
Signal:
Filling Factor
Q-value
Conversion Factor:
Notes: ----- Meeting Notes (9/8/16 15:18) -----
finite-element modeling gives me these solutions
----- Meeting Notes (3/30/17 16:58) -----
mention TE011
square root sign
cavity resonators

## Slide 4
Motivation
Hydrogenase
Understanding the mechanism of biological hydrogen conversion
2H+ + 2e- ⇌ H2
Extremely Sample Limited
Resonator design for Nano-liter volumes at X-band
Materials Research
Samples of Limited Availability
Systems on a Chip/Micro-fluidics
Protein Single Crystals

## Slide 5
Single Protein Crystal [FeFe] Hydrogenase
D. desulfuricans HydAB
Lubitz, Ogata, Rüdiger, Reijerse, Hydrogenases, Chem. Rev., 114, 2014
0.5-50 nanoliter
To date no single crystal studies have been performed on [FeFe] Hydrogenase
Using FTIR, EPR, NMR, Raman, and NRVS a convincing catalytic cycle has been hypothesized.
However, it is very difficult to hypothesize the F-clusters and g-tensors axis, protein single crystals are needed.
Notes: ----- Meeting Notes (3/30/17 16:58) -----
7-10 A

## Slide 6
Increasing Sensitivity for Limited Samples

## Slide 7
Planar Micro-Resonators at X-band
Resonator fabricated using standard lithography techniques
Loop of 0.5 mm diameter or smaller
Resonance is achieved using strip-line impedance
Sapphire substrate has a Q-value of 180 with a high filling factor
A factor of 2 is realized over the Bruker MD5 dielectric resonator for a sample size of 0.3 x 0.2 x 0.2 mm3
Narkowicz, Suter, Stonies, J. Magn. Reson., 175, 2005

## Slide 8
Micro-Helix at X-band
1.2 mm Tall
5 windings 9.7 GHz
0.15 mm silver wire
0.4 mm inner diameter
Higher concentration of B1
Higher Q-value of 360 due to surface vs. volume currents
  More homogeneous field
    Easier placement of small samples

## Slide 9
Micro-Helix Preliminary Results
Simulation | DR ε=10 | PMR 0.5 mm | Helix 0.4 mm
Freq. [GHz] | 9.67 | 9.58 | 9.41
Q0-value | 15000 | 186 | 360
Norm. Ss | 1 | 2.09 | 9.5
Cmax [mT/W1/2] | 0.65 | 3.7 | 3.5
Cave [mT/W1/2] | 0.65 | 2.2 | 3.3
Simulated
Measured
LiPC Crystal 0.3 x 0.2 x 0.2 mm3
Norm. Signal Saturable
MD5: 1
PMR: 2.2
Micro-Helix: 10.2
Sidabras, Reijerse, Suter, Lubitz, In Preparation
Notes: Add slide numbers
----- Meeting Notes (9/8/16 15:29) -----
LiF SHow which is which
----- Meeting Notes (3/30/17 16:58) -----
micro helix advantages
scale on

## Slide 10
Micro-Helix Preliminary Results
Experimental Results:
Tyrosine (YD*) Radical in Photosystem II
Experiment performed at 120 K.
0.3 x 1012 spins/G
Single Crystal
0.3 x 0.3 x 0.25 mm3
Empty Resonator
Experimental Results:
Mn2+ signal
Sidabras, Reijerse, Suter, Lubitz, In Preparation
Notes: ----- Meeting Notes (3/30/17 16:58) -----
measurements doing on a crystal
0.3 x 0.3 x 0.5

## Slide 11
244 GHz Single-Mode Resonators
TE311 has z-component of electric field. This leads to significant microwave leakage through modulation slots.
40%
2.5%
Bucket Resonator is simplest structure
Sacrifices Q-value and B1 distribution for sample volume.
A TE011 mode improves by least a factor of 2

## Slide 12
244 GHz Re-Entrant Semi-Spherical TE011
Geometry | Cylindrical TE011 D/L=1 | Re-Entrant Semi-Spherical TE011
Frequency | 240.5 | 243.6
Q0-value | 4300 | 1200
Signal, Su | 0.96 | 2.15
Signal, Ss | 0.04 | 0.06
Cmax [mT/W1/2] | 3.4 | 3.8
B0
B0
0.4 mm

## Slide 13
Motivation

## Slide 14
Uniform Field Resonators
Mett, Froncisz, Hyde, Rev. Sci. Instrum, 72(11), 2001
Anderson, Mett, Hyde, Rev. Sci. Instrum., 73(8), 2002
Hyde, Mett, Anderson, Rev. Sci. Instrum. 73(11) 2002
Mett, Sidabras, Hyde, Appl.. Magn. Reson., 31, 2007
Design resonator with a region-of-interest at cut-off
Cut-off waveguide has a wavelength of infinity (uniform)
An end-section is needed to properly match the cut-off region
By transferring a short to an open (1/4 wavelength) you can properly match an infinite wavelength.

## Slide 15
Uniform Field Resonators at W-band
Sidabras, Mett, Sarna, Hyde, J. Magn. Reson., Under Review

## Slide 16
Re-Entrant Uniform Field TE01U at Q-band
Higher concentration of B1
Lower Q-value for the same EPR signal
  Significantly more homogeneous field
Geometry | Cylindrical TE011 D/L=1 | UF TE01U | UF TE01U Re-Entrant
Frequency | 34.3 | 34.5 | 34.0
Q0-value | 13000 | 5900 | 1880
Signal, Su | 1686.2 | 1232.1 | 1782.4
Signal, Ss | 158.8 | 132.2 | 188.0
Λmax [mT/W1/2] | 1.06 | 0.65 | 0.94
Λave [mT/W1/2] | 0.52 | 0.52 | 0.83
% Diff | 52 | 19.2 | 12.2
Sample Geo. | 1.8 ID x 2.8 OD | 1.8 ID x 2.8 OD | 1.8 ID x 2.8 OD

## Slide 17
Re-Entrant Uniform Field TE01U at Q-band

## Slide 18
3D Printed Resonators at Q-band
Current 3D printing technology may be good enough for Q-band tolerances
Solid brass or silver is available using molding techniques
Higher precision techniques exist for ABS plastics that can be plated.

## Slide 19: Acknowledgements
Prof. Dieter Suter
Dr. Nick Cox
Dr. Hideaki Ogata
Dr. Anton Savitsky
Dr. Edward Reijerse
Prof. Wolfgang Lubitz
Marie Skłodowska-Curie Individual Fellowship
http://act-epr.org
Notes: ----- Meeting Notes (3/30/17 16:36) -----
it is very difficult to figure out the F-clusters, g tensors axis were located in NiFe and help develop the calculations also g tensor axis for DFT!!
no real finalized DFT and g-tensor interactions
need g-tensor axis system for EDNMR/ENDOR
case study example
extremely limited introduction slide
in general, biological systems are of low symmetry so crystal structures don't give g-tensor info
in single crystals you can measure the g/A at room temp, and no averaging out the tensors
spings/G/s W-band 10^(7or8)
how improtant the structural axis interaction is, low symmetry high complex system
----- Meeting Notes (3/30/17 16:58) -----
mention Blanck PMR resonators
g strain higher at high frequency.
Q-band with Blank is in the works