## Slide 1: Resonator Development for Studying Protein Single Crystals of Limited Dimensions at X-band
Jason W. Sidabras
D. Suter, E.J. Reijerse, A. Savitsky, W. Lubitz

## Slide 2
Motivation
Extremely Sample Limited
Nano-liter volumes at X-band
Materials Research
Samples of Limited Availability
Systems on a Chip/Micro-fluidics
Protein Single Crystals
Hydrogenase
Understanding the mechanism of biological hydrogen conversion
2H+ + 2e- ⇌ H2

## Slide 3: Why Single Protein Crystals?
Lubitz, Ogata, Rüdiger, Reijerse, Chem. Rev., 114, 2014
Foerster, Stein, Brecht, Ogata, Higuchi, Lubitz, JACS, 125, 2003
50-1000 nanoliter
Active Site
EPR with Single Crystals:
Hydrogenase catalytic cycle has various redox states and several are paramagnetic
Disentangle angular dependencies of g- and A-tensors
Give insights into molecular interactions at the active site by relating to the X-ray Crystal Structure
Quantum Chemical Calculations to understand catalytic mechanism the full g-tensors axis. EPR data helps develop the calculations
Notes: ----- Meeting Notes (9/8/16 15:12) -----
Talk less about it. Just that its been done before. Move to the front.
Orientation dependence.
----- Meeting Notes (3/30/17 16:58) -----
Protein single crystal
disentagle
full tensor information, best for even testing DFT
Quantum Chemical Calcuations not DFT

## Slide 4: Why Single Protein Crystals?
Lubitz, Ogata, Rüdiger, Reijerse, Chem. Rev., 114, 2014
Foerster, Stein, Brecht, Ogata, Higuchi, Lubitz, JACS, 125, 2003
50-1000 nanoliter
Active Site
EPR with Single Crystals:
Hydrogenase catalytic cycle has various redox states and several are paramagnetic
Disentangle angular dependencies of g- and A-tensors
Give insights into molecular interactions at the active site by relating to the X-ray Crystal Structure
Quantum Chemical Calculations to understand catalytic mechanism the full g-tensors axis. EPR data helps develop the calculations
Notes: ----- Meeting Notes (9/8/16 15:12) -----
Talk less about it. Just that its been done before. Move to the front.
Orientation dependence.
----- Meeting Notes (3/30/17 16:58) -----
Protein single crystal
disentagle
full tensor information, best for even testing DFT
Quantum Chemical Calcuations not DFT

## Slide 5
Single Protein Crystal [FeFe] Hydrogenase
Lubitz, Ogata, Rüdiger, Reijerse, Hydrogenases, Chem. Rev., 114, 2014
0.5-50 nanoliter
To date no single crystal studies have been performed on [FeFe] Hydrogenase
Using FTIR, EPR, NMR, Raman, and NRVS a convincing catalytic cycle has been hypothesized.
However, it is very difficult to hypothesize the F-clusters and g-tensors axis, protein single crystals are needed.
D. desulfuricans HydAB
Notes: ----- Meeting Notes (3/30/17 16:58) -----
7-10 A

## Slide 6
Finite-Element Modeling
Ansys HFSS solves the full-wave Maxwell’s Equations for a given geometry and boundary condition
Build and Test on the Bench
Characterize the resonators and compare to simulation
Measure coupling, frequency, and Q-value
Experiments with Standard Samples
Further characterize the resonators using EPR
Sample handling (very important)
Resonator Introduction
Rectangular TE102
Magnetic field magnitude
Electric field magnitude
Notes: ----- Meeting Notes (3/30/17 16:58) -----
title methods

## Slide 7
Feher, Sensitivity Considerations in Microwave Paramagnetic Resonance Absorption Techniques, 36(2). 449,1957
Hyde, Froncisz, Advanced EPR: Applications in Biology and Biochemistry, Elsevier, 1989, Ch. 7: Loop-gap resonators, 277
Based on a critically coupled resonator and reflection bridge configuration
Resonator Characteristics
Cylindrical TE011
Signal:
Filling Factor
Q-value
Conversion Factor
Notes: ----- Meeting Notes (9/8/16 15:18) -----
finite-element modeling gives me these solutions
----- Meeting Notes (3/30/17 16:58) -----
mention TE011
square root sign
cavity resonators

## Slide 8
Increasing Sensitivity for Limited Samples

## Slide 9
Tiny (< 500 μm) loop
Distributed capacitance
    approx. λ/2 length
    High dielectric substrate
Very High Conversion Factor
Q-value < 100 for Rogers material
Q-value approx. 200 for sapphire substrate
Narkowicz, Suter, Stonies, Planar microresonators for EPR experiments, J. Magn. Reson., 175, 2005
Planar Micro-Resonators for EPR

## Slide 10
Planar Micro-Resonators Results
Simulation | MD5 DR ε=10 | 0.5 mm PMR
Freq. [GHz] | 9.67 | 9.58
Q0-value | 15000 | 186
Norm. Su | 1 | 2.09
Cmax [mT/W1/2] | 0.65 | 3.7
Cave [mT/W1/2] | 0.65 | 2.2
All simulations are performed with a 0.3 x 0.3 x 0.3 mm3 sample at room temperature
Experimental Results:
Tyrosine (YD*) Radical in Photosystem II
Single Crystal 0.3 x 0.3 x 0.5 mm3
0.6 x 1012 spins/G
Notes: ----- Meeting Notes (3/30/17 16:58) -----
measurements doing on a crystal
0.3 x 0.3 x 0.5

## Slide 11
Micro Helix at X-band
Higher concentration of B1
Higher Q-value (surface vs. volume)
  More homogeneous field
    Easier placement of small samples
1.2 mm Tall (0.15 mm wire)
5-6 windings 9.7 GHz

## Slide 12
Micro-Coils in NMR
Size < 0.5 mm inner diameter
Parallel Plate Capacitance to resonate
Q-value < 30
Good Sensitivity enhancement
is realized for limited samples
K. Yamauchi, J.W.G. Janssen, A.P.M. Kentgens, Implementing solenoid microcoils for wide-line solid-state NMR, J. Magn. Reson. 167 (2004)
Hans Janssen, Andreas Brinkmann, Ernst R. H. van Eck, P. Jan M. van Bentum, Arno P. M. Kentgens, Microcoil High-Resolution Magic
Angle Spinning NMR Spectroscopy, JACS, 128, 2006
Notes: ----- Meeting Notes (9/8/16 15:29) -----
Very low Q.
900 MHz systems

## Slide 13
Micro-Helix Preliminary Results
Simulation | MD5 DR ε=10 | PMR 0.5 mm | Helix 0.4 mm
Freq. [GHz] | 9.67 | 9.58 | 9.41
Q0-value | 15000 | 186 | 360
Norm. Su | 1 | 2.09 | 9.5
Cmax [mT/W1/2] | 0.65 | 3.7 | 3.5
Cave [mT/W1/2] | 0.65 | 2.2 | 3.3
Simulated
Measured
LiPC Crystal 0.3 x 0.2 x 0.2 mm3
MD5: 1
PMR: 2.2
Micro-Helix: 10.2
Notes: Add slide numbers
----- Meeting Notes (9/8/16 15:29) -----
LiF SHow which is which
----- Meeting Notes (3/30/17 16:58) -----
micro helix advantages
scale on

## Slide 14
Micro-Helix Preliminary Results
Experimental Results:
Tyrosine (YD*) Radical in Photosystem II
Experiment performed at 120 K.
0.3 x 1012 spins/G
Single Crystal
0.3 x 0.3 x 0.25 mm3
Empty Resonator
Experimental Results:
Mn2+ signal
Notes: ----- Meeting Notes (3/30/17 16:58) -----
measurements doing on a crystal
0.3 x 0.3 x 0.5

## Slide 15
Coupling with Large Mutual Inductance
Coupling constant k is much larger than needed for maximum power transfer.
Instead the coupling constant is optimized for maximum impedance transfer.
The primary is then matched to the load for typical power transfer.
Output voltage is enhanced and noise is reduced.
Sidabras, Mett, Hyde, MRI surface-coil pair with strong inductive coupling, Rev. Sci. Instrum., 87, 124704, 2016

## Slide 16
Coupling with Large Mutual Inductance
New Coupling PC Board:
Loop Resonant at 17 GHz
Planar loop reduces E-field
Coupling implemented
Current Coupling Scheme:
Grounded loop has a very high self-resonance (~30 GHz)
Currently coupling the anti-parallel mode
  Significant capacitance lowers modes and mutual inductance
Parallel mode is measured at ~2 GHz
~1.7x coupling enhancement
(17x total over MD5)

## Slide 17: Acknowledgements
Future Work
Sample handling, including a goniometer
Repeatable Micro-Helix fabrication and manufacturing techniques
Test Helix below 50 K
Prof. Dieter Suter
Dr. Nick Cox
Dr. Hideaki Ogata
Dr. Anton Savitsky
Dr. Edward Reijerse
Prof. Wolfgang Lubitz
Marie Skłodowska-Curie Individual Fellowship
http://act-epr.org

## Slide 19
Coupling with Large Mutual Inductance At 400 MHz
Theoretical Enhancement of 3 (spiral) * 3 (coupling)
13 mm
Sidabras, Mett, Hyde, MRI surface-coil pair with strong inductive coupling, Rev. Sci. Instrum., 87, 124704, 2016